SCOPE 9 ЛИК

Sofia, 07.03.2024

Scientific, Research, Testing and Calibration Laboratory for Lighting Technique

at Technical University, Sofia

Management and laboratory address:

1797 Sofia, 8 St. Kliment Ohridski Blvd.

To perform testing of:

| Type of the scope: flexible** |  |  |  |
| --- | --- | --- | --- |
| № | Tested products | Type of test / characteristic | Testing methods (standard /validated method) |
| 1 | 2 | 3 | 4 |
| 1 | 1. LIGHT SOURCES 1.1. Incandescent lamps 1.2. Halogen filament lamps 1.3. Mercury gas discharge lamps 1.4. Sodium gas discharge lamps - low and high pressure 1.5. Metal halogen discharge lamps 1.6. Car Lamps 1.7. Luminescent lamps - tubular and compact 1.8. Different others. types of lamps (such as compact LED lamps, tube light-emitting diodes, light-emitting retrofit) 1.9. Special lamps | Power | БДС EN 13032-1, cl. 5.3, table 2 БДС EN 13032-4, cl. 4.3 БДС EN 60809 cl. 4.8 and Annex С CIE Publication № 121 |
| 2 |  | Electrical current | БДС EN 13032-1, cl. 5.3, table 2 БДС EN 13032-4, cl. 4.3 БДС EN 60809 cl. 4.8 and Annex С CIE Publication № 121 |
| 3 |  | Voltage | БДС EN 13032-1, cl. 5.3, table 2 БДС EN 13032-4, cl. 4.3 БДС EN 60809 cl. 4.8 and Annex С CIE Publication № 121 |
| 4 |  | Luminous flux | БДС EN 13032-1, cl. 5.5, cl. 6.1.2, cl. 6.1.3, table 3 БДС EN 13032-4, cl. 6.2 БДС EN 60809, Annex С |
| 5 |  | Light intensity | БДС EN 13032-1, cl. 5.4 and table 2 БДС EN 13032-4, cl. 6.2 |
| 6 |  | Brightness, overall brightness | БДС EN 13032-1, cl. 5.6 and table 2 БДС EN 13032-4, cl. 6.7 |
| 7 |  | Resistance to increased electromotive voltage | БДС ЕN 60432-1, cl. 2.9, Annex Е |
| 8 |  | Strength of bolt fastening | БДС EN 60432-1, cl.2.5 and tables 3 a nd 4; Annex С БДС EN 60968, cl. 9 |
| 9 |  | Electrical insulation resistance of the socket | БДС EN 60432-1, cl. 2.6; Annex А БДС EN 61195, cl. 2.4 |
| 10 |  | Lamp heating | БДС EN 60432-1, cl. 2.4, table 2, Annex К БДС EN 60968, cl. 10 БДС EN 60969, Annex А БДС EN 61195, Annex В БДС EN 60360 |
| 11 |  | Flashing time | БДС EN 60188, cl. 1.4, Annex А and В БДС EN 61167, cl. 4.5 БДС EN 60969, Annex А |
| 12 |  | Coordinates of color | БДС EN 60188, cl. 1.4.7, Annexes В and С БДС EN 61167, cl. 4.5 БДС 8.882, repealed but not replaced for requirements and methods БДС EN ISO 11664-3 CIE Publication № 15 |
| 13 |  | Light Efficiency (Energy Efficiency Class) | БДС EN 50285, cl. 6 Regulation (EC) № 244, Annex III |
| 14 | 1.10. Light-emitting diodes | Luminous flux | БДС EN 13032-1, cl. 5.5, cl. 6.1.2, cl.6.1.3, table 3 CIE Publication № 127 |
| 15 |  | Light intensity | БДС EN 13032-1, table 2 CIE Publication № 127 |
| 16 |  | Coordinates of color | БДС EN ISO 11664-3 CIE Publication № 15 БДС 8.882, repealed but not replaced for requirements and methods CIE Publication № 127 |
| 17 |  | Work temperature | CIE Publication № 127 |
| 18 | 2. LIGHTERS (including LED luminaries): 2.1. Fixed lights for general lighting 2.2. Built-in luminaires 2.3. Street and road lighting fixtures 2.4. Portable general purpose luminaires 2.5. Projectors 2.6. Lamps with incandescent lamps incorporating transformers or converters 2.7. Portable lighting for parks 2.8. Hand-held luminaires 2.9. Portable luminaries, attractive for children 2.10. Photo and cinema lamps (non-professional) 2.11. Stage lighting, TV movie studios and cinema studios (indoor and outdoor installation) 2.12. Illuminators for emergency lighting 2.13. Light garlands 2.14. Ultra low voltage lighting systems with incandescent lamps 2.15. Luminaires with limitation of temperatures on the surface of the housing 2.16. Lighting for use in hospital and health care facilities 2.17. Lighting fixtures 2.18. Special lighting fixtures 2.19. Lighting for motor vehicles | Electrical insulation resistance | БДС EN 60598-1, cl. 10.2 |
| 19 |  | Electrical insulation strength | БДС EN 60598-1, cl. 10.2 |
| 20 |  | Light distribution | БДС EN 13032-1, cl. 5.4, cl. 5.7, cl. 5.8, cl. 6, cl. 7, cl. 8, tables 2, 3 and 4, Annexes А, В БДС EN 13032-4, cl. 6.5 |
| 21 |  | Light intensity | БДС EN 13032-1, cl. 5.4 and table 2 БДС EN 13032-4, cl. 6.5 |
| 22 |  | Efficiency coefficient (without light-emitting diodes luminaires) | БДС EN 13032-1, cl. 5.5 |
| 23 |  | Heating of various items | БДС EN 60598-1, cl. 12.4 and tables 12.1 and 12.2 |
| 24 |  | Temperature of the supporting surface | БДС EN 60598-1, cl. 12.4 and tables 12.1 and 12.2 |
| 25 |  | Mechanical strength | БДС EN 60598-1, cl. 4.13.1, 4.13.2; table 4.3 |
| 26 |  | Mechanical strength of the hanger | БДС EN 60598-1, cl.4.14.1, cl. 4.14.2 |
| 27 |  | Roll resistance | БДС EN 60598-1, cl. 4.4.14 |
| 28 |  | Length of insulating distances on the surface of the insulation and through the air | БДС EN 60598-1, cl. 11.2, table 11.1 |
| 29 |  | Protection against direct contact | БДС EN 60598-1, cl. 8.2.1, cl. 8.2.5 БДС EN 60529 |
| 30 |  | Protection against indirect contact | БДС EN 60598-1, cl. 7.2.1, cl. 7.2.8, cl. 7.2.11 |
| 31 |  | Ability to place on combustible material | БДС EN 60598-1, cl. 12.5 and table 12.3 |
| 32 |  | Illumination | БДС EN 12464-1,2, cl. 6.1 |
| 33 |  | Overall brightness | БДС EN 13032-1, cl. 5.6 and table 2 БДС EN 13032-4, cl. 6.7 |
| 34 |  | Electricity | БДС EN 13032-1, cl. 5.3, table 2 БДС EN 13032-4, cl. 4.3 |
| 35 |  | Power | БДС EN 13032-1, cl. 5.3, table 2 БДС EN 13032-4, cl. 4.3 |
| 36 | 3. LIGHTING INSTALLATIONS 3.1. For internal natural and artificial lighting. 3.2. For street lighting. 3.3. For sports lighting. 3.4. For artistic and architectural lighting. 3.5. For tunnel lighting. 3.6. For emergency lighting. | Illumination | MI 40 (repealed but not replaced), item 4 БДС EN 12464-1, 2 БДС EN 13201-4 CR 14380, cl.9 and Annex А БДС EN 1838, Annex А БДС EN 12193, cl. 6 |
| 37 |  | Luminance | MI 40 (repealed but not replaced), item 4 БДС EN 12464-1, 2 БДС EN 13201-4 CR 14380, cl. 9 and Annex А БДС EN 1838, Annex А БДС EN 12193, cl. 6 |
| 38 |  | Illumination uniformity | MI 40 (repealed but not replaced), item 4 БДС EN 12464-1, 2 БДС EN 13201-4 CR 14380, cl.9 and Annex А БДС EN 1838, Annex А БДС EN 12193, cl. 6 |
| 39 |  | Luminance uniformity | MI 40 (repealed but not replaced), item 4 БДС EN 12464-1, 2 БДС EN 13201-4 CR 14380, cl. 9 and Annex А БДС EN 1838, Annex А БДС EN 12193, cl. 6 |
| 40 | 4. SIGNALING MEANS 4.1. Traffic lights 4.2. Signal lights, 4.3. Headlights 4.4. Light - signaling devices, etc. light marking | Light intensity | БДС EN 12368, cl. 8.2; БДС EN 12352, cl. 6.2; Ordinance № 14, Annex № 9 БДС 4704 cl. 3.2.1, cl. 3.2.2, cl. 5.1.1, cl. 5.1.2 БДС 10916, from cl. 4.3 to cl. 4 |
| 41 |  | Luminance uniformity of the illuminated field | БДС EN 12368, cl. 8.3 БДС EN 12352, cl. 6.3 |
| 42 |  | Light intensity coefficient (Specific reflection coefficient) | БДС EN 12368, cl. 8.4 БДС EN 12352, cl. 6.2 |
| 43 |  | Coordinates of color | БДС EN 12368, cl. 8.5/8.6 БДС EN 12352, cl. 6.4 Ordinance № 14/2000, Annex № 1 to art. 101, para. 3 UNECE Regulation № 43 /Rev. 2/ 1995, Annex 3, item 9.4 БДС 4704 (repealed but not replaced), cl. 5.1.3 БДС 10916, cl. 4.2 |
| 44 |  | Luminance factor е | БДС EN 12368, cl. 8.5/8,6 БДС EN 12352, cl. 6.4 |
| 45 |  | Skip Rate (Coefficient of light oversight) | ECE Regulation № 43 /Rev. 2/ 1995, Annex 3, item 9 БДС 9973– cl. 4.16 |
| 46 | 4.5. Road signs: 4.5.1. Fixed, fixed vertically; 4.5.2. With internal lighting; 4.5.3. Reflective bars; 4.5.4. Variable message signs 4.5.5. Reflective road studs. | Coordinates of color | БДС EN 12899-1, cl. 4 and cl.7, Annex А БДС EN 12966-1, cl. 5.5 БДС EN 1463-1, Annexes А, В and С |
| 47 |  | Luminance factor е | БДС EN 12899-1, cl. 4 and cl. 7, Annex А БДС EN 12966-1, cl. 5.5 БДС EN 1463-1, Annexes А, В and С |
| 48 |  | Light intensity coefficient (Specific reflection coefficient) | БДС EN 12899-1, cl. 4 and cl.7, Annex А БДС EN 12966-1, cl. 5.5 БДС EN 1463-1, Annexes А, В and С |
| 49 |  | Luminance uniformity | БДС EN 12899-1, cl. 4 and cl.7, Annex А БДС EN 12966-1, cl. 5.5 БДС EN 1463-1, Annexes А, В and С |
| 50 |  | Luminance, Luminance ratio | БДС EN 12966-1, cl. 5.5 |
| 51 | 4.6. Road markings (road marking paint) | Coordinates of color | БДС EN 1436, cl. 4.4, table 6 and 7; Annex C |
| 52 |  | Luminance factor е | БДС EN 1436, 4.4, table 6 and 7; Annex C |
| 53 |  | Light intensity coefficient (Specific reflection coefficient) | БДС EN 1436–cl. 4.2 and cl. 4.3, Annexes А and В |
| 54 | 4.7. Signs for vehicle registration numbers and trailers 4.8. Identifying signs for moving vehicles | Coordinates of color | БДС ISO 7591, cl. 7 UNECE Regulation № 69 UNECE Regulation № 70 UNECE Regulation № 104, item 7.1 and Annex 6 |
| 55 |  | Luminance factor е | БДС ISO 7591, cl. 7 UNECE Regulation № 69 UNECE Regulation № 70 UNECE Regulation № 104, item 7.1 and Annex 6 |
| 56 |  | Light intensity coefficient (Specific reflection coefficient) | БДС ISO 7591, cl. 7 UNECE Regulation № 69 UNECE Regulation № 70 UNECE Regulation № 104, item 7.1 and Annexes 1 and 7 |
| 57 | 4.9. Stop-signal triangle | Coordinates of color | UNECE Regulation № 27, item 2 and Annex 5 |
| 58 |  | Luminance factor е | UNECE Regulation № 27, item 2 and Annex 5 |
| 59 |  | Light intensity coefficient (Specific reflection coefficient) | UNECE Regulation № 27, item 4 and Annexes 5 and 6 |
| 60 |  | Unifirmity of Light intensity coefficient | UNECE Regulation № 27, item 4 and Annex 5 |

References:

| 1 | *БДС 8.882:1984 (repealed but not replaced) | Colorimetry. Basic methods for measuring color |
| --- | --- | --- |
| 2 | *БДС 4704:1984 (repealed but not replaced) | Vehicle electrical equipment. Headlights. |
| 3 | CIE Publication № 127:2007 - Publication of the International Commission on Illumination (CIE) | Measurement of LEDs. |
| 4 | CIE Publication № 15:2018 - Publication of the International Commission on Illumination (CIE) | Colorimetry. |
| 5 | CIE Publication № 121:1996 - Publication of the International Commission on Illumination (CIE) | Photometry and goniophotometry of luminaries. |
| 6 | Regulation (EC) № 244:2009, Annex III | Regulation implementing Directive 2005/32/ EC of the European Parliament and of the Council with regard to ecodesign requirements for non-directional household lamps |
| 7 | MI (Methodological Instructions) 40-85-CCNTP - Headquarters Standardization: 1985 | Methods for measurement and evaluation of artificial lighting. |
| 8 | СД CR 14380, (CEN report): 2005 | Applied lighting. Tunnel Lighting |
| 9 | UNECE Regulation № 27/Addendum 26:1998 | Technical requirements for the approval of a stop-signal triangle. |
| 10 | UNECE Regulation № 43/Addendum 42:1995 | Technical requirements for wheeled vehicles and parts thereof. |
| 11 | UNECE Regulation № 69/Addendum 68:1995 | Uniform Conditions for Approval of Rear Bodies for Slow Vehicles and their Trailers. |
| 12 | UNECE Regulation № 70/Addendum 69:1987 | Uniform Conditions for Approval of Rear Bodies for Long and Long Vehicles. |
| 13 | UNECE Regulation № 104/Addendum 103:1998 | Uniform Conditions for Approval of Light-reflecting Markings for Slow-Driven Vehicles and Their Trailers. |
| 14 | Ordinance № 14/2012, SG, № 86 and № 87 | Ordinance on airports and airport insurance (Main characteristics of airports). |

**Flexible scope: Implementing a new version of standards/documents or standards/documents replacing them is allowed. An updated list of standards/documents and their dated versions is provided by laboratory.

To perform calibration of:

| Type of the scope: fixed |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| № | Measuring instrument | Measured value, unit of measure | Measurement Range | Measurement uncertainty | Calibration method |
| 1 | 2 | 3 | 4 | 5 | 6 |
| 1 | Luxmeter | Illumination, lx | from 50 lx to 1000 lx | 2,5 % (к=2) | МЕ 11-13/2021 (РПК 17:2021) |

Reference: ME 11-13/2021 (РПК 17:2021) Methodology for Calibration of Luxmeters.

Note: The calibration activity is only performed in the laboratory.